Features:

| - Oscilloscope + multimeter + waveform generator 3 in 1 - 80MHz analog bandwidth 250MSa/s sampling rate - Afterglow selection - AI waveform preview - Self-calibration function | - Support DC 15W fast charge (come with the fast charger) - Support PC software analysis - 2.8-inch 320*240 TFT color LCD display - 6000 counts true RMS large capacitance multimeter - 3 trigger modes: auto mode, normal mode, single mode |
| --- | --- |

﻿﻿Specification:

| General characteristics |  |
| --- | --- |
| Power supply | AC 100V〜240V、50Hz〜60Hz;DC 15W fast charge |
| Power consumption of the whole machine | Oscilloscope range≤400mA, multimeter range≤200mA |
| Battery | 18650 lithium battery 2000mAh x 2 |

Oscilloscope

| Number of channels | Channel CH1, Channel CH2 |
| --- | --- |
| Analog bandwidth | 80MHz (single channel), 50MHz (dual channel) |
| Maximum real-time sampling rate | 250MSa/s (single channel)）,125MSa/s (dual channel) |
| Waveform interpolation | (Sin x) /x |
| Record length | When single channel, the maximum is 6K; when dual channel, the maximum is 3K |
| Horizontal scale range | 5nS/div〜50S/div, 1 > 2、5steps |
| AD converter | 8-bit resolution, simultaneous sampling of each channel |
| Vertical scale range | 1 X: 10mV/div-10V/div; 10X: 100mV/div-100V/div |
| Vertical measurement range | ±4 grid |
| Optional analog bandwidth limit | 20MHz |
| Low Frequency Response (-3db) | ≤10Hz at BNC |
| Vertical Gain Accuracy | ±3% in "Normal" or "Average" acquisition mode |
| Input resistance | 1MΩ/15 |
| Input Coupling | DC, AC, GND |
| Probe attenuation | 1X, 10X |
| Support probe attenuation factor | 1X, 10X, 100X |
| Input protection voltage | 300Vpp |
| Trigger Type | Edge Trigger |
| Trigger Mode | Auto, Single, Normal |
| Trigger level | ±4div |
| Edge-triggered slope | Rising edge, falling edge, rising edge or falling edge |
| Voltage parameter measurement | peak-to-peak, RMS, average, peak, maximum, minimum |
| Time parameter measurement | frequency, period, duty cycle, positive pulse width, negative pulse width |
| Cursors | Voltage difference between cursors, time difference between cursors Smart cursor 1 group, the capture point keeps tracking the current waveform when the cursor moves, and the current cursor position information, voltage difference and time difference are displayed below. |
| Spectral Analysis FFT | Yes |
| Storage capacity | 2000 groups of DMM data, 200 groups of waveform records, with storage file manager |
| Filtering | high pass, low pass |
| Math operation | FFT |
| Afterglow option | Yes (afterglow effect is required to make Lissajous graphics) Afterglow selection is divided into multi-level optional. |
| Computer real-time communication | USB communication |
| Waveform zoom | Yes |
| AI waveform preview | a small window is added above the main interface to display the thumbnail of the waveform for preview, and with an indicator to indicate that the current main interface waveform is at a certain position in the entire collected waveform. |

Multimeter

| Function | Range Uncertainty | Uncertainty |
| --- | --- | --- |
| DC voltage | 600.0mV〜1000V | 0.75%rdg + 10dgt |
| AC voltage | 600.0mV〜750V (TRMS,frequency range：40Hz〜1KHz) | 1.0%rdg + 10dgt |
| DC current | 60.00mA/500.0mA, 6.000A/10.00A Overload protection: fuse type fuse 500mA/250V,10A/250V | 1.2%rdg + 10dgt |
| AC current | 60.00mA/500.0mA, 6.000A/10.00A (TRMS) Overload protection: fuse type fuse 500mA/250V,10A/250V | 1.5%rdg + 10dgt |
| Resistance | 600.0Ω～60.00M | 1.0%rdg + 5dgt |
| Capacitance | 10.00nF〜1.000mF 1.000mF〜100.0mF | 2.0%rdg + 10dgt 5.0%rdg + 10dgt |
| Frequency | 10.00Hz〜10.00MHz | 1.0%rdg + 5dgt |
| Diode test | Open circuit voltage is about 3V,maximum test current is about2mA |  |
| On-off test | ≤50Ω beep，≥600Ω display 0L |  |

Waveform Generator

| Waveform frequency | Sine wave: 1Hz～5MHz Square waves: 1Hz～500KHz Triangular wave: 1Hz～500KHz |
| --- | --- |
| Amplitude | Fixed-point output, maximum 5Vpp (50Ω)——0Vpp, 0.1Vpp, 0.2Vpp, 0.5Vpp, 1Vpp, 2Vpp, 3Vpp, 5Vpp |
| Frequency Resolution | 0.2Hz |
| Vertical resolution | 8 bit |
| Output Impedance | 50Ω |
| Square waves pulse width accuracy | ±20nS |